# Introduction
A Versatile Labtool
jithinbp.in/vLabtool
Videos, sample code
User Manual
A cost effective, multi-pronged data acquisition and control instrument to turn your computer into a workbench for science and electronics experiments .
## Plug and Play multitool with Experiment Oriented Apps.

Version 12.04 or Greater
Version 2.7 or Greater

## Feature Summary
Calibrated Against Professional equipment for high accuracy
Feature	Description	Socket Labels	Range
Analog Inputs	Up to 10 analog inputs with various voltage ranges. / Programmable gain and offset control. / 12-bit voltage measurement. Up-to 1MSPS sampling. / 5ppm, 3.3V Voltage Reference	CH1, CH2 / CH3, / CH4 - CH7 / SEN / I2V	±16Volts / ±3.3Volts / 0 − 3.3V / 0 − 3.3V / ±3.3mA
Analog Outputs	Programmable Voltage Sources (12-bit) / Programmable Constant Current Source	PVS1 / PVS2 / PVS3 / PCS	±5V / ±3.3V / 0 − 3.3V / 0 − 3.3mA
Digital Inputs	4 × TTL compatible digital Inputs. / Logic Analyzer / Frequency Counter (0-30MHz), / time measurement routines / ±2.5ppm, 12M Hz Reference oscillator	ID1 - ID4	0-5 V
Waveform / Generators	Sine/triangle wave. / Programmable frequency : 28-bit resolution / Manual Amplitude Control / Sine wave / Fixed Frequency / Fixed Amplitude / 4 × Phase Correlated Square Waves / Frequency / Phase difference maximum resolution / Duty Cycle maximum resolution / Also functions as simple state selectable output	Wavegen / SINE / SQ1 - SQ4	5Hz - 2MHz / 0-4 Volts / 14 kHz / ±4Volts / 0 − 5Volts / 10Hz-16MHz / 15nS / 15nS
Wireless Nodes	Battery Powered add-on units. / Power Source / Acts as a wireless bridge for various sensors / Multiple Nodes can be used together / Can also control relays and read voltages	Cell/Adapter	3.7-4.2 V
Data Buses	I2C : Master / Commonly Used by numerous sensor ICs / SPI : Master / Chip Select pins / UART	SCL,SDA / SCK,SDI,SDO / CS1 , CS2
Expansion Slot	20 - pin Socket designed to accommodate / add-on modules. / Selection includes SPI,I2C,Analog IO,Digital IO, / 16MHz TTL output, and a DC Sine Wave.

# Installing the software
.
The device has been tested on Debian based distributions: Ubuntu 12.04 and above, and Raspbian Wheezy. Windows support pending.
## System Requirements
Processor	Entry level PCs as well as Netbooks running Intel Atom processors are / supported. No special hardware required. Tested with the Raspberry Pi 2..
Display	Graphical Utilities are scalable down to 800x600 without scrollbars entering / into the picture. / Accelerated Graphics are only required if responsive 3D plots are being used

## Installating dependencies
Run the following commands on a terminal (CTRL-ALT-T) :
1. sudo apt-get install python-qt4 python-opengl python-qt4-gl python-scipy
  python-numpy ipython-qtconsole python-setuptools python-sip
2. sudo easy_install pyqtgraph
3. sudo easy_install vLabtool
## List of Installed Software
The Following is an incomplete list of graphical applications that have been installed along with the Python module, and can be run from a terminal
1. vLabtool-scope, 2-channel Oscilloscope, 4-chan Logic Analyzer, Peripheral control and curve fitting abilities.
2. vLabtool-stream, Designed to continuously plot values from any function from the vLabtool Python module that returns an integer/decimal.
3. vLabtool-sensorDemo, Auto-Detects sensors connected to the I2C expansion slot, and allows plotting their readings in real time., This requires the sensor class to be installed, and several are supported by default.
4. vLabtool-console, Launches an iPython console with the initialization functions already carried out., A plotting window is also incorporated, and can used with a single function call.
5. vLabtool-wirelessDemo, Auto Detects wireless nodes in the vicinity, and lists their addresses and connected sensors, It plots data from multiple sensors in real time.
6. vLabtool-bodePlots, Uses the programmable sine wave generator and analog inputs to obtain the frequency and phase response of filters.
7. vLabtool-diodeIV, Plots the Current-Voltage characteristics of a diode. Also forms a surface plot of multiple curves to study the temperature dependence
# Testing : Oscilloscope Application
Copy the following command and run it on a terminal: vLabtool-scope

Screenshot of the oscilloscope utility plotting a sine waves generated by the waveform generators.
The data has also been fitted, and the results displayed in the bottom section.
## Control widgets in the Analog Signals section
	CH1 : Channel 1 of the oscilloscope
Remappable	Not restricted to ’CH1’ input socket. This is a very flexible acquisition channel, and can be mapped to any of the analog inputs via a drop down menu located at the top corner, It also Features 12-bit offset control via a slider.
	CH2 : Channel 2 of the oscilloscope
Fixed	This channel can only record the signals input to the socket labeled CH2.
	TRIGGER : Allows selecting the channel to trigger the oscilloscope
Level	The rotary dial sets the trigger level, and this is also reflected in a marker / located on the plot.
Channel	The drop down menu allows selecting the trigger channel, If CH1 is selected, / the trigger signal will be obtained from the channel it has been mapped to.
	Time Gap : Specify the time scale of the oscilloscope
Dial	Set the delay(in µs) between each successive data point.

	Sine Fit :Uses Scipy to fit the data against a sine function
Menu 1	Select input data channel for fitting, Uses Scipy.optimize and obtains frequency, amplitude, phase and offset values.
Menu 2	Select a second input data channel for fitting.
checkbox	Overlay the fitted results.
Results	The fitted values: Amplitude, Freq, Phase, Offset are displayed in the message window at the bottom of the screen.
	Lissajous : XY Plots
Menu 1	Select input data channel for X-axis.
Menu 2	Select input data channel for Y-axis.
show	Enable/Disable Lissajous overlay.
Animate	Pause data acquisition and animate the formation of the X-Y plot.

## Control widgets in the Peripherals section

Screenshot Of the oscilloscope utility showing the peripherals section. The Lissajous plot has been enabled in the Analog section.
PVS1	Range : ±5V , Set the output voltage on PVS1 socket., Max Current 5mA.
PVS2	Range : ±3.3V , Set the output voltage on PVS2 socket., Max Current 5mA.
PVS3	Range : 0 − 3V , Set the output voltage on PVS3 socket., Max Current 5mA.
PCS	Range : 0 − 3.3mA, Set the output current on PCS socket. Current subject to load resistance, PCS and PVS3 are linked, and only one can be used at a time.

Wavegen	Range : 0 − 2M Hz, Set the output frequency of the waveform generator., Max resolution 0.04Hz., Amplitude control(±1mV − ±4V ) via physical knob located on the device.
	Phase Correlated Square Waves (SQ1,SQ2,SQ3,SQ4)
Frequency	Range : 0 − 8M Hz, Set the output frequency of all four square waves., / Resolution:15nS.
Output	Select the Output socket to configure.
Phase Duty	Set the phase difference (0 − 100%) , between the selected waveform and SQ1.
Cycle	Set the Duty Cycle (0 − 100%) , of the selected waveform.

## Control widgets in the Digital section

Screenshot of the digital signals section in the four channel logic analyzer mode. Four PWM signals from SQ1 - SQ4 have been connected to Digital inputs ID1-ID4. CH3 has been set to record only rising edges, while CH4 records only falling edges
Active Channels	Number of inputs to digitize. Single channel and two channel modes can record / up to 60 seconds delay between consecutive level changes with a resolution of / 15nS. Three and Four channel modes can only record up to 1 ms delay between / consecutive level changes.
Acquisition time	Time to sample inputs before fetching results. The LA continuously refreshes / data at this interval..

	Single Channel Options
Channel	The input to record. Select from ID1,ID2,ID3,ID4.
Edges	Types of Level changes to record. Select from options such as Falling edges only, / rising edges only, all edges.
	Trigger
Channel	Enable level change trigger on selected channel.
Edges	Type of level change to use as the trigger.
	3/4 channel modes
CH1-CH4	Set the type of level changes to record for each channel when the logic / analyzer is in either the 3 channel, or the 4 channel mode.
	Measure Time Intervals > Important! Freeze logic analyzer before using this.
Edge 1	Select input channel and edge type that will start the timer.
Edge 2	Select input channel and edge type that will stop the timer.
Go	Run the measurement.
Results	Results will be displayed here in seconds units.

# Testing : Wireless Units Demo
Copy the following command and run it on a terminal: vLabtool-wirelessDemo
## Connecting a new wireless node carrying some sensor
Screenshot of the wirelessDemo application
1. Plug in a sensor on to a wireless node by matching the pins labeled Vdd, GND, SCL, SDA. These pins appear in the same order on most sensor modules, and no extra wires should be needed.
2. Make sure that the option Register New Nodes on the Application is selected.
3. Plug In the Battery onto the wireless node.
4. Observe that the software has auto detected your wireless node as well as found out the address of the sensor you connected.
5. Unselect Register New Nodes.
  If you wish, you may move the cursor over a sensor’s address and the software will guess it’s model number/type.
6. Click on , Refresh Node List, This loads a few controls for each wireless node, and also a drop down selection of the connected sensors.
## Logging Data from a sensor
After Following the steps in the previous section, Select the appropriate sensor and click on Go, The software will continuously plot values fetched from the sensor.
## Changing sensor parameters
Open the auto-generated menu at the bottom left corner, and change parameters. A few examples include gain, sampling rates, and data-type selection.
## Testing : Data logger Application Demo
This application plots the return values from any of the functions defined in the Python Module as long as they return integers or decimals.
Copy the following command and run it on a terminal: vLabtool-stream
## Choosing a command and starting the logger
Screenshot of the streaming application.
1. Refer to the programmer’s manual for a list of valid commands, and pick the function you would like to monitor, The manual can be opened from the help menu.
2. Paste the command name with correct arguments into the application. And now click on Monitor. The app includes a few common commands and appropriate arguments packed into a drop down menu linked to the command entry box.
## Additional Functionality Accessible through menus
Voltmeter	Load Widgets that perform a single voltage measurement per click.
Amplifiers	Load a widget to set the gain on the voltage input channels.
Wavegen	Load a slider that allows setting the frequency of the sine wave output.
IV sources	Load widgets that, enable setting programmable voltage and current.
Timing	Load time measurement widgets, These include duty cycle and frequency.
Console	Insert an iPython Console with init commands preloaded.
help	Help menu that allows access to the programmer’s manual as well as experiment specific HTML files.

# Pin Definitions and Functions
## Analog Inputs - Measure Voltages with 12-bit resolution
CH1	Range: ± 16V, Gain: up to 32×, Separate AC coupled input.
CH2	Range: ± 16V, Gain: up to 32×, Separate AC coupled input.
CH3	Range: ± 3.3V, Gain: up to 32×, 12-bit Offset control. / *manual gain control by grounding Rg via resistor R. g=1+{10K} over {R_g}
CH4	Range: 0 – 3.3V, Gain: up to 32×, 12-bit Offset control.
CH5	Range: 0 – 3.3V, Gain: up to 32×, 12-bit Offset control.
CH6	Range: 0 – 3.3V, Gain: up to 32×, 12-bit Offset control.
CH7	Range: 0 – 3.3V, Gain: up to 32×, 12-bit Offset control.
SEN	Range: 0 – 3.3V, Internally Pulled up to 3.3V via 5.1KOhm resistor.
AN2	Range: 0 – 3.3V, Available on the expansion slot.
AN3	Range: 0 – 3.3V, Available on the expansion slot.
	Internally Monitored
5 V	The supply voltage, Measures the power supply voltage of the USB port.
V+	Measures the output voltage. / Monitor this to verify if external loads are affecting it

## Analog Outputs
	Waveform Generators
WG1	0-2MHz Sine/ Triangle Waveform Generator, Range: ±4V. Manual Amplitude Control via dial located on the side. / 28- bit (0.04Hz) resolution
SINE	Fixed 14KHz Sine Waveform Generator, Range: ±4V.
	Programmable Voltage and Current Sources (12-bit)
PVS1	Range: ±5V , Up to 10mA.
PVS2	Range: ±3.3V , Up to 10mA.
PVS3	Range: 0 − 3.3V , Up to 10mA.
PCS	Range: 0 − 3.3mA. / Subject to Load resistance. Up to 3.3V voltage drop

## Digital Inputs - Measure Logic Levels with 15nS resolution
	Logic Analyzer, Frequency Counter, Timing Functions
ID1	Range: 0 − 5V, 0-32MHz.
ID2	Range: 0 − 5V, 0-32MHz.
ID3	Range: 0 − 5V, 0-32MHz.
ID4	Range: 0 − 5V, 0-32MHz.

## Digital Outputs
	4× PWM output with 15nS frequency, phase, and DCycle control
SQ1	Range: 0 − 5V.
SQ2	Range: 0 − 5V.
SQ3	Range: 0 − 5V.
SQ4	Range: 0 − 5V.

## Data Buses - I2C , SPI, UART
	I2C (Inter-Integrated Circuit) Master: Up to 4MHz Clock speed
Application	Tested With Various sensors, MPU6050, MLX90614, HMC5883L, BMP180, / TSL2561, SHT21 etc. / These measure angular velocity, acceleration, temperature, humidity, luminosity, pressure, and magnetic fields.
Vdd	3.3V Power output.
GND	Power return path.
SCL	I2C Clock Output.
SDA	I2C Data Line.
	SPI(Serial-Peripheral Interface) Master: Available on the 20-pin Expansion Slot. Supports all four modes of operation.
SCK	SPI Clock Output, Tested Up To 16MHz.
SDO	SPI Data Output (MOSI).
SDI	SPI Data Input (MISO).
O1	Chip Select 1.
O2	Chip Select 2.

## A Few Definitions For The Expansion slot
DDS	16MHz , 0-3.3V TTL output. This is also the reference clock for the 0-2MHz sine/triangle wave generator, and can be used to create phase correlated wavegen add-ons.
AN2	Analog Input, Range: 0 − 3.3V.
AN3	Analog Input, Range: 0 − 3.3V.

## Additional Nomenclature
GND	Short for GROUND. This is the reference voltage, and is taken as 0 volts. It is also connected to the ground level of the power source (laptop/tablet/RPi etc.)
V+/V-	Power Outputs, ±9V . Current limit 20mA per channel.

# Wireless Nodes

## On-board Transceiver : 2-way radio link
Capable of communicating with multiple wireless nodes to control and receive sensor data payloads
## Wireless Subunits : 3-Byte addresses

I2C	Data bus for Interfacing with commonly available sensors. Measure physical parameters with minimal mechanical interference.
SPI	Another data bus used by some devices. Of interest to the vLabtool are precise waveform generators, and high resolution analog instruments.
RGBout	Control WS2812B arrays , Daisy chained RGB LEDs with 16 million shades each.